## SOFTWARE FMEA
ID | Function under assessment | Potential Failure Mode | Cause of Failure | Effect of Failure | Occurrence | Severity | Detection | RPN | Control | Occurence | Severity | Detection | RPN
R2 | Software | Display issues mean that toolkit content or functionality does not display correctly | Browser side incompatibilities (eg browser, JS, OS versions) | Features of the toolkit do not display correctly or user cannot properly interact with the toolkit as features do not work as intended. Consequently clinician is not advised of appropriate actions. Clinician will use own clinical judgement to decide on course of action. Potential for increased likelihood of clinical error as a result. | 2 | 3 | 2 | 12 | Rigorous testing for browser compatibility (Tactuum) RDS parent app provides statement within of whch browsers are compatible with toolkit content. Link to Tactuum browser and accessibility statements can be repeated within toolkit "About" section. | 1 | 3 | 3 | 9
R3 | Software | Output is not noticed or ignored. | Design and formatting of page content. | Clinician does not see critical clinical information. Will resort to own knowledge base with increased likelihood of clinical error. | 2 | 3 | 2 | 12 | Usability testing will evaluate that key information is clear and visible User education/orientation sessions to ensure good clinical awareness of how to use the app. | 1 | 3 | 1 | 3
R4 | Release process | Toolkit content or functionality are not up to date. | Failure in release process to properly push out updates following a change to the content. | Clinician will use out of date content and own clinical judgement to decide on course of action. Increased likelihood of clinical error as a result. | 2 | 3 | 1 | 6 | Updates are automatic on web version of app. On mobile version, updates are automatically delivered as soon as the user has an Internet connection. In current version of Umbraco, user needs to manually accept updates; in new platform updating will be automatic. | 1 | 3 | 1 | 3
R5 | Content | Content is no longer aligned with up to date evidence/ best pratice. | Failure to keep up to date with clinical evidence | Toolkit no longer reflects the current clinical guidance. In consequence, recommended diagnosis/treatment may be out of date and patient does not get optimal care. | 2 | 3 | 2 | 12 | Regular clinical reviews of evidence. Process and responsibilities in place for ensuring that this is converted into guidance, approved and translated into the toolkit in a timely manner. Post-market surveillance activities | 1 | 1 | 2 | 2
R7 | Software | Toolkit cannot be accessed. | For web-based application - poor Internet connectivity or loss of Internet connectivity. | Toolkit unavailable /features of the toolkit do not display correctly or user cannot properly interact with the toolkit as features do not work as intended. Clinician will use own clinical knowledge to decide on course of action. Potential for increased likelihood of clinical error as a result. | 2 | 3 | 1 | 6 | Encourage clinicians to download mobile app as it will continue to operate in case of Internet connectivity failure. RDS contract with Tactuum means that failure in Internet availability is classified as Priority 1 - Critical/Urgent, response within 4h of notification via Support Portal, best endeavours to address within 8h. | 1 | 3 | 1 | 3
R8 | Static Content | Content has errors, either in clinical content or spelling and grammatical errors | Content not properly reviewed | Lack of clarity in static content and guidance or clinical error in guidance given. Potential for error in recommended diagnosis/treatment for the patient, leading to suboptimal care. | 3 | 1 | 2 | 6 | 1. Proof reading of content for grammar and spelling 2. Clinical review of content as part of usability testing. | 1 | 1 | 2 | 2
R9 | Software | Toolkit could be used by a member of the public | Toolkit is publicly available on internet and can be accessed by anyone | Members of the public may not be able to properly interpret the information and may use the toolkit as a substitute for professional advice | 3 | 1 | 1 | 3 | Add statement to toolkit (e.g. in About-Disclaimer section) that it should not be used by qualified health and social care staff and should not be used by members of the public as a substitute for professional advice | 1 | 1 | 1 | 1
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0

## Risk Acceptability
Criteria for Risk Acceptability
Severity
Value | Description | Criteria
1 | Negligible | Inconvenience or temporary discomfort.
3 | Minor | Results in temporary injury or impairment not requiring professional medical intervention
5 | Serious | Results in injury or impairment requiring professional medical intervention
7 | Catastrophic | Results in permanent impairment, life threatening injury or death
|  | Occurrence |  |  |  |  | Detection
Value | Description | Criteria |  |  |  | Value | Description | Criteria
1 | Improbable | Unlikely to happen |  |  |  | 1 | Frequent | Likely to be detected
2 | Occasional | Can happen but not frequently |  |  |  | 2 | Occasional | Can happen but not frequently
3 | Frequent | Likely to happen |  |  |  | 3 | Improbable | Unlikely to be detected
Description | Criteria | [6 empty] | Threshold | Upper Value
High | Risk must be reduced or eliminated | [6 empty] | 21 | 72
Medium | Risk should be Reduced as Far as Possible; accept only if risk has been reduced as far as possible | [6 empty] | 9 | 20
Low | Further risk reduction not required | [6 empty] | 1 | 8
Occurrence x Detection
Severity | Value | 1 | 2 | 3 | 4 | 6 | 9
| 1 | Low | Low | Low | Low | Low | Medium
| 3 | Low | Low | Medium | Medium | Medium | High
| 5 | Low | Medium | Medium | Medium | High | High
| 7 | Low | Medium | High | High | High | High